## Slide 1: WP3.2.1: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
Evelyn Freney, Olivier Favez, Valerie Gros, Jean Sciare, Andre Prevot, Jay Slowik, Francesco Canonaco
Deliverables
To provide homogeneous quality-controlled ACSM datasets at a European scale
To homogenize calibration and data processing and in particular the source apportionment of organic aerosol data.
To assure the precision and accuracy of these measurements through comparison with various co-located instruments e.g TEOM-FDMS, SMPS, OC-EC etc.
With the help of
Bologna, Italy. October 2016

## Slide 2
1. Homogenous quality-controlled ACSM datasets at a European scale
  Intercomparison campaign took place March/April 2016. A total of 22 instruments within the ACTRIS network participated.
  In order to accommodate all applications to the intercomparison exercise, two separate calibration exercises were organised.
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
Site | Intercomparison Q-ACSM ToF-ACSM | 
Hohenpeissenberg |  | 
FMI –Hyytiala |  | 
Cyprus |  | 
Melpitz |  | 
Finokalia |  | 
London |  | 
Bologna |  | 
Madrid |  | 
Vavihill |  | 
SMEAR (UHEL) |  | 
Zurich |  | 
JFJ |  | 
Bucharest |  | 
Barcelona |  | 
PUY |  | 
SIRTA |  | 
CapCorse |  | 
Tartu |  | 
ToFwerk |  | 
Calib: Quad –ACSM
Calib: ToF-ACSM
DNA
Quad-ACSM: 4th to 14th of March
(15 instruments)
ToF-ACSM: 21st of March to 18th of April
(12 instruments (6 ToF-ACSM))
2016 ACTRIS-2 ACSM intercomparison exercise at the ACMCC

## Slide 3
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
Quad-ACSM
arrival
3rd March
4th to 7 th of March
Ambient measurements
7th to 10 th of March
Calibration,
IE NO3, IE SO4
11th to 14 th of March
Ambient measurements
Phase 2:
ToF-ACSM
Research Intercomparison
21st March
acmcc@lsce.ipsl.fr
21st to 24th of March
Ambient measurements
24th to 27th of March
Calibration/PAM experiments
27th of March to 18th of April
Ambient measurements
  Organisation:
  Three to four days pre-calibration intercomparison.
  Calibration using single and mixed inorganic solutions
  Three to four days post-calibration intercomparison (extended for ToF-ACSM).
2. To homogenize calibration and …..
2016 ACTRIS-2 ACSM intercomparison exercise at the ACMCC

## Slide 4
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
1. Instrument falls within ± 30% of the « reference » instrument
2. Instrument performance evaluated against the Median using the Z-score method
PM1 SIRTA (REF)
  Organic, sulphate, nitrate, Ammonia
  PM1 of all instrument compared with PM1 of SIRTA instruments.
  Dotted line indicates ± 30% error
  light points represent TEOM-FDMS, black points represent ACSM data
More discussion of the intercomparison results and evaluation
Thursday 09:30-11:15: WP8 (A. Wiedensohler) European Centre for Aerosol Calibration
3. Precision and accuracy of these measurements were evaluated with various co-located instruments from such as TEOM-FDMS, SMPS, OC-EC Sunset Field analyzer, PILS-IC, filter sampling …).
2016 ACTRIS-2 ACSM intercomparison exercice at the ACMCC

## Slide 5
  Deliverables
  21/25 of the ACTRIS ACSM instruments were successfully evaluated.
  New operating and calibration procedures were put in place.
  Instruments were compared with external reference instruments as well as with the SIRTA reference instrument.
  Workshop with the ACTRIS community for training/discussion and improvement
  Actions:
  Calibration reports need to be distributed to each of the instruments PIs.
  New SOPs need to be implemented
  Announcement of new inter-laboratory exercise (Oct 2018 )
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
Work in process
Introduction of new calibration procedures into SOP
Autumn 2017:
For new instruments or those unable to attend first intercomparison,
A smaller calibration exercise will be annonced.
2016 ACTRIS-2 ACSM intercomparison exercise at the ACMCC

## Slide 6
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
2a. Data submission protocol:
Current data submission format is outlined on the NILU page

## Slide 7
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
Current format is outlined on the NILU page
Common template for all ACSM (Quad/ToF)/(HR/C/-ToF/Quad) and AMS instruments
The measurement principle of all these instruments remains the same.
Metdata allows for instrument type to be disquished.
Particle concentrations are expressed as hourly averaged g m-3
In order to quantify the atmospheric variability on an hourly basis, the standard
deviation is also reported alongside the hourly average.
Missing data is represented as -99999
2a. Data submission protocol:

## Slide 8
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
2a. Data submission protocol:
What is EETA?
 Utility for exporting ACSM/AMS data to external databases
Output files use EBAS format
Handles reporting of “standard” outputs and PMF matrices
Distributed through ACTRIS-ACSM mailing list.
Current version: 2.6
 Q-ACSM fully supported
Partial support for other instruments
Still a few stability issues, but no known calculation problems
ACTIONS
Goal of “final” version by mid-2017
Support for all types of aerosol mass spectrometry instruments
Source apportionment analysis

## Slide 9
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
ACTIONS
To increase the amount of data being submitted from ACSM instruments to the database.
NRT submission of level 0 data.
2a. Data submission protocol:

## Slide 10
Canonaco et al., 2013
Current state
Current SoFi version
user-friendly
fast (exploits multiprocessors for PMF calculations and during data analysis)
smart storage using HDF technology
Email to: francesco.canonaco@psi.ch
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
2a. Data processing and source apportionment  SoFi – The interface in Igor Wavemetrics Pro
Notes: 10

## Slide 11
Actions:
finalize/test automatic selection of the sources
Test on other data sets
build/test on-line source apportionment (Automated Rolling  AuRo SoFi)
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor
2a.Data processing and source apportionment  SoFi – The interface in Igor Wavemetrics Pro
Notes: 11

## Slide 12
  ACTIONS:
  1. ACMCC:
  Calibration reports need to be distributed to each of the instruments PIs. (Dec 2016)
  Announcement of new inter-laboratory exercise
  - basic calibration exercise 2017 (new and/or uncalibrated instruments),
  - larger (research) intercomparison 2018
  2. ACMCC + PSI : Implementation of existing and new SOPs
    In current SOPs calibrations are recommended every two months. We aim to request groups to send their calibration results every 6 months as a means to follow instrument performance.
  All: Data submission
    Increase the amount of submitted data. Set up of NRT of level0 data
Task 3.2.1
WP3: Development of a measurement standardization and data submission protocol for aerosol particle mass spectroscopy based on Aerosol Chemical Speciation Monitor