## Slide 1: WP3.2.3: Development of a measurement standardization and a data submission protocol for VOCs based on PTRMS and TOF-MS measurements
Truman Wright and Rupert Holzinger, Utrecht University , the Netherlands
Sebastien Dusanter and Stephan Sauvage, Ecole Nationale Superieure des Mines de Douai;
Anja Claude, DWD;
Stephan Reimann, EMPA

## Slide 2: Content
1. Current state of the art (without standard operation protocol, SOP)
  1.1 PTR-TOF-MS in a nut-shell
  1.2 The kinetic method and its issues
  1.3 The calibration method
2. The SOP (so far…)
  2.1 Recommended settings
  2.2 Merging of the two methods for the mutual benefit
  2.3 In-situ gas-standard additions for data quality assurance (QA)
3. Quality control (QC) tools (exploiting the gas-standard additions)
  3.1 PTR parameters (O2+/H3O+, E/N, sample flow, sample humidity)
  3.2 TOF-MS parameters (transmission, peak shape, baseline signal)
4. QA/QC protocol
5. Next steps

## Slide 3: 1.1 PTR-TOF-MS in a nut-shell
…
Proton transfer reactions (PTR):
H3O+ + R  H2O + RH+
… occur at the collisional rate for almost all VOCs.
Hansel et al., 1995:

## Slide 4: 1.2 The kinetic method and its issues
Issues
H3O+ not the only primary ion, also ~10-50% water dimer (H2O H3O+),
transmission through the mass spectrometer depends on ion mass,
uncertain reaction rate constant, k, (20-30%),
fragmentation for some compounds.
[R]=
Holzinger et al., 2010:
Hansel et al., 1995:
[R]=
[R] concentration of VOC [molec cm-3]
VMRR volume mixing ratio of compound [mol/mol]
N number concentration of gas in drift tube [molec cm-3]
t reaction time (residence time in drift tube) [s]
Primary ions:
[H3O+] concentration [ions cm-3]
I19 measured intensity [counts per second]
k19 reaction rate constant with H3O+ [cm3s −1molecule−1 ]
Product ions:
[RH+] concentration [ions cm-3]
IRH+ measured intensity [counts per second]
Water dimer (H2OH3O+, second primary ion):
[H2OH3O+] concentration [ions cm-3]
I37 measured intensity [counts per second]
k37 reaction rate constant with H2OH3O+ [cm3s −1molecule−1 ]
=
VMRR

## Slide 5: 1.3 The calibration method
De Gouw and Warneke, 2007:
S determined through frequent in-situ gas standard additions:
“Measure NIRH+ for a known volume mixing ratio.”
NIRH+ normalized measured intensity [normalized counts]
XR unitless compound specific humidity factor (experimentally determined)
S compound specific sensitivity
[normalized counts per ppb]
Advantages/disadvantages:
Straightforward
Accurate for calibrated compounds
No worries about transmission, fragmentation, and rate constant
Humidity factor remains an issue
Uncalibrated signals are often not reported or very inaccurate

## Slide 6: 2.1 Recommended settings
Compromise between
  low H2O H3O+/H3O+ ,
  fragmentation of VOCs,
  and O2+/H3O+.

## Slide 7: 2.2 Merging of the two methods for the mutual benefit
The benefit:
S can be easily measured with a gas standard.
Physical parameters (E/N, k19, k37, transmission, and XR ) can be retrieved from simple measurements of S, humidity, and fragmentation.

## Slide 8: 2.3 In-situ gas-standard additions for data quality assurance (QA)
16 compounds
Covered mass range 33-373 Da
Some compounds fragment
Range of XR values
No overlapping signals

## Slide 9: 3.1 Quality control: PTR parameters
Ratio O2+/H3O+
Extraction to TOF region

## Slide 10: 3.1 Quality control: PTR parameters
Flow into PTR-MS
??
Inlet
flow

## Slide 11: 3.1 Quality control: PTR parameters
…
E/N (Td) = a(frx)2 + b(frx) + c (frx given as a percentage (%)) |  |  |  | 
m/z of parent compound | m/z of fragment | A | b | c
137.132 | 81.070 | 0.00587 | 0.210 | 82.7
223.064 | 207.032 | 0.0417 | -1.37 | 114.6
371.101 | 355.070 | 0.00110 | 0.0641 | 97.1
E/N

## Slide 12: 3.1 Quality control: PTR parameters
Sample humidity & H2O leak from ion source
Humidity factor, Xr

## Slide 13: 3.2 Quality control: TOF-MS parameters
Baseline signal
R. Holzinger: PTRwid: A new widget-tool for processing PTR-TOF-MS data
Peak shape
http://www.atmos-meas-tech.net/8/3903/2015/amt-8-3903-2015.pdf

## Slide 14: 3.2 Quality control: TOF-MS parameters
Transmission

## Slide 15: 4. QA/QC protocol
Preliminary Protocol:
30s: FH2O low, dry air
30s: FH2O low, humid air
30s: FH2O normal, humid air
40s: Finlet=20sccm, 2 Std injections
40s: Finlet=40sccm, 2 Std injections
40s: Finlet=60sccm, 2 Std injections
40s: FH2O low, dry air, E/N=100Td, 2 Std injections
40s: FH2O normal, dry air, E/N=100Td, 2 Std injections
40s: FH2O high, dry air, E/N=100Td, 2 Std injections
Retrievals:
Inlet flow (4,5,6)
O2+/H3O+ (3)
Extraction to TOF good? (1,2)
H2O leak from ion source (2,3)
E/N (4,5,6)
Transmission (4 thr 9)
Sensitivity, S (4,5,6)
XR (7,8,9)
Baseline signal (1 thr 9)
Peak shape (1 thr 9)
General output:
accuracy, precision, inconsistency alerts

## Slide 16: 5. Next steps
Get the community involved
Construct a stand-alone standard addition device that can be shipped to PTR-MS users for testing our SOP and QA/QC protocols.
Develop a simple measurement protocol for full QA (in ~5 minutes in automated systems)
Data submission protocol
Publications