## Slide 1
Development of reference instruments for multi-wavelength absorption measurements
Thomas Müller and Alfred Wiedensohler
Leibniz Institute for Tropospheric Research, Leipzig
ACTRIS-2 WP3 Workshop, Bologna, October 2016

## Slide 2
Method and combined uncertainty
absorption = extinction - scattering
 | extinction | scattering
Instrument(s) | Three CAPS PMex at 447, 530 and 655 nm | Three wavelength Nephelometer Aurora4000, 450, 525, 635 nm
Adjustment of wavelengths | Interpolation rule based on Ångström exponents | 
Contributions to combined uncertainty of absorption | Noise Baseline drift Gas absorption (ambient trace gases) Path length calibration Non-linearity Temperature stability | Noise Baseline drift Calibration factor Truncation Temperature stability
Corrections: Except ambient gas absorption (O3, NO2,..), all effects contribution to the combined uncertainty can be corrected. Corrections for minimizing base line drift are currently under development (cf. page 4). |  | 

## Slide 3
Calibration and Quality assurance
Calibration of Scattering measurement
Zero adjust
CO2 span calibration
Calibration of Extinction measurement
Quality assurance
(checks and calibrations)
Daily: Nephelometer zero
Weekly: Nephelometer span and CAPS’s path lengths
Single scattering albedos higher 0.9 require more frequent checks. A clear recommendation is not yet possible.
Calibration of CAPS’s effective pathlengths and non linearity.
Weekly: Check for noise levels and baseline drift.
Comparing light scattering and extinction for non absorbing aerosols.

## Slide 4
CAPS baseline drift
Extinction and baseline signal for particle free air
‘Jumps’ in baseline signal (> 0.15 Mm-1)
Obviously, previous baseline measurement induced an irreversible change of baseline signal
Baseline drift < 0.15 Mm-1 is compensated by instrument
Methods for corrections and flagging are under development
prev.
baseline
meas.

## Slide 5
Non-linearity of CAPS
Non-linearity was determined several times in the last 18 months.
No significant change was observed.

## Slide 6
Combined uncertainty for 1 minute averages
Theoretically, longer averaging times would reduce instrumental noise and uncertainty.
But, baseline drift on time scales of 1 to 15 minutes limits the benefit of longer averaging times.
Analysis of baseline drift and development of methods
correction and flagging is ongoing.