## Slide 1: WP8: Calibration facilities, Aerosol Chemical Monitor Calibration Centre Evaluation protocol
Evelyn Freney, Olivier Favez, Valerie Gros, Jean Sciare
Yunjiang ZHANG, Tanguy AMODEO
Bologna, Italy. October 2016
It’s complicated!

## Slide 2
It’s complicated! : some examples
ACSM measurement artefacts
Crenn et al., 2015 + Frohlich et al., 2015
Observed for the first time after the ACTRIS-1 ACSM intercomparison (2013)
Discrepancies within organic mass fractions from one instrument to another
Seemed to be linked to NO3 contents

## Slide 3
It’s complicated! : some examples
ACSM measurement artefacts
Recent study conducted on the various PSI AMS and ACSM’s
Can be observed during direct injection of ammonium nitrate (in acquisition mode)
Published Sept. 2016

## Slide 4
It’s complicated! : some examples
ACSM measurement artefacts
Recent study conducted on the various PSI AMS and ACSM’s
Can be observed during direct injection of ammonium nitrate (in acquisition mode)
Published Sept. 2016
2013 intercomparison
ACTRIS-2 2016 intercomparison
ACTRIS-2: this relationship appears to be dependant on vapourizer dirtiness

## Slide 5
ACSM measurement artefacts
Can we define harmonized correction protocols (when needed) ?
Need for a better « calibration » of organic mass spectra
Secondary organic aerosol generated from know compounds and from different UV exposures
Organic fragments show significant variation between instruments - especially for m/z 44.
A second « research » thematic that was adressed during the 2016 intercomparison
It’s complicated! : some examples

## Slide 6
Quad-ACSM
arrival
4th to 7 th of March
Ambient measurements
11th to 14 th of March
Ambient measurements
7th to 10 th of March
Calibration,
IE NO3, IE SO4
Phase 2:
ToF-ACSM
21st to 24th of March
Ambient measurements
24th to 27th of March
Calibration/PAM experiments
27th March to 18th April
Ambient measurements
Calibration set-up
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 7
PM1 SIRTA (REF)
  PM1 of all instrument compared with PM1 of SIRTA instruments.
  Dotted line indicates ± 30% error
  light points represent TEOM-FDMS, black points represent ACSM data
2016 ACTRIS-2 ACSM intercomparison exercice at the ACMCC: Evaluation of instruments preformance
2. Instrument performance evaluated against the Median using the Z-score method
  Organic, sulphate, nitrate, Ammonia
1. Instrument falls within ± 30% of the « reference » instrument
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 8
1. Instrument falls within ± 30% of the « reference » instrument
2. Instrument performance evaluated against the Median using the Z-score method
PM1 SIRTA (REF)
  Organic, sulphate, nitrate, Ammonia
  PM1 of all instrument compared with PM1 of SIRTA instruments.
  Dotted line indicates ± 30% error
  light points represent TEOM-FDMS, black points represent ACSM data
2016 ACTRIS-2 ACSM intercomparison exercice at the ACMCC: Evaluation of instruments preformance
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 9
Validation of reference instrument
SIRTA instrument was identified as reference instrument
Instrument has been sampling at SIRTA since 2011. The longest existing ACSM data set.
It is regularily calibrated and maintained by experienced SIRTA staff.
Comparison with external reference methods
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 10
Evalulation criteria (List of some of potential errors)
ACMCC: ACTRIS2_First intercomparison analysis and reporting
Calibration error
Transmission efficiency
Collection efficiency
Detection limit (Zero measurements)
Instrument artefacts

## Slide 11
Evalulation criteria (Error associated with calibration)
As expected, high variability within response factors from one instrument to the next → Each instrument has is own response factor
But, less expected, for a given instrument:
Calibration repeatability: Variation in calibration values (up to 3 calibrations were performed on each instrument)
Discrepancies between average values from user’s reported calibration and those obtained at the ACMCC:
Deviation = 24%
Same person using same calibration set-up at ACMCC:
Deviation = up to 14% , partly due to different particle concentration reading directly from CPC vs. from the ACSM instrument
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 12
Evalulation criteria (Error associated with instrument transmission and collection)
TE: Transmission efficiency.. The Lens TE varies depending on the particle size measured (25%).
Error in TE
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 13
Evalulation criteria (Error associated with instrument transmission and collection)
CE: Correction for collection efficiency of indiviual instruments depends on composition and phase
Instruments are then corrected using a chemical dependant collection efficiency (CDCE).
Assuming that when particles contain a high fraction of ammonium nitrate that they are more efficiency sampled (more liquid)
The instrument to instrument variation in the CDCE =5%
Error in TE
Error in CE
TE: Transmission efficiency.. The Lens TE varies depending on the particle size measured (25%).
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 14
Evalulation criteria (Conclusion)
Calibration repeatability : Variation in calibration value (3 to 4 calibrations were performed on each instrument)
Error = 14%
TE: Transmission efficiency error.. Depending on the particle size measured the lens TE varies..
(25%)
CE: Collection efficiency correction.
Depending on the particle chemical composition and morphology..
Error in CDCE based on chemical composition
5%
+
29%
Propagation of errors
Error1
Error2
Error3
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 15
Evalulation criteria (Conclusion)
Calibration repeatability : Variation in calibration value (3 to 4 calibrations were performed on each instrument)
Deviation = 14%
TE: Transmission efficiency. Depending on the particle size measured the lens TE varies.. (Deviation = 25%)
CE: Collection efficiency correction.
Depending on the particle chemical composition and morphology.. (30% from literature)
Error in CDCE based on chemical composition
5%
+
29%
Propagation of errors
** This is a first estimate of instrument errors.
ACTIONS
TE: Use longterm ILC data to determine TE error on ambient data (comparing with external FIDAS instrument).
Species by species uncertainties will be determined for all instruments.
Instrument detection limits will be detected from zero measurements.
Species by species uncertainties (Crenn et al., 2016).
Error1
Error2
Error3
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 16
Method 2
Instrument performance evaluated against the Median using the Z-score method (ISO 13528 (2005))
  Organic, sulphate, nitrate, Ammonia
Average concentration of instrument i
Median concentration
Standard deviation
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 17
Pre-calibration intercomparison. Comparison with reference instrument
Internal instrument reference
This can be compared with values of the instrument prior to the instrument coming to the centre
Some examples of individual instrument comparisons
BEFORE
Serial # | RFNO3 | RIENH4 STD | RIESO4 STD | RIESO4 MIX | RIENH4 MIX
Original | 5.20e-11 | 9.57 | 0.54 | - | -
Calibrated | 6.24e-11 | 5.38 | 0.32 | 0.70 | 3.50
*** All settings did not load up correctly when instrument arrived at station
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 18
Post-calibration intercomparison. Comparison with reference instrument
Internal instrument reference
This can be compared with values of the instrument after the instrument goes back to the centre
Some examples of individual instrument comparisons
AFTER
Serial # | RFNO3 | RIENH4 STD | RIESO4 STD | RIESO4 MIX | RIENH4 MIX
Original | 5.20e-11 | 9.57 | 0.54 | - | -
Calibrated | 6.24e-11 | 5.38 | 0.32 | 0.70 | 3.50
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 19
Internal instrument reference
This can be compared with values of the instrument after the instrument goes back to the centre
Pre-calibration intercomparison. Comparison with reference instrument
Some examples of individual instrument comparisons
BEFORE
Serial # | RFNO3 | RIENH4 STD | RIESO4 STD | RIESO4 MIX | RIENH4 MIX
Original | 2.03e-11 | 6.23 | 0.42 | - | -
Calibrated | 4.28e-11 | 4.81 | 0.32 | 0.98 | 5.54
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 20
Internal instrument reference
This can be compared with values of the instrument after the instrument goes back to the centre
Pre-calibration intercomparison. Comparison with reference instrument
Some examples of individual instrument comparisons
Serial # | RFNO3 | RIENH4 STD | RIESO4 STD | RIESO4 MIX | RIENH4 MIX
Original | 2.03e-11 | 6.23 | 0.42 | - | -
Calibrated | 4.28e-11 | 4.81 | 0.32 | 0.98 | 5.54
AFTER
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 21
Pre-calibration intercomparison. Comparison with reference instrument
Internal instrument reference
This can be compared with values of the instrument after the instrument goes back to the centre
Some examples of individual instrument comparisons
BEFORE
Serial # | RFNO3 | RIENH4 STD | RIESO4 STD | RIESO4 MIX | RIENH4 MIX
Original | 5.63e-12 | 5.11 | 0.76 | - | -
Calibrated | 3.21e-11 | 5.98 | 0.4 | 0.58 | 5.1
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 22
Internal instrument reference
This can be compared with values of the instrument after the instrument goes back to the centre
Post-calibration intercomparison. Comparison with reference instrument
Some examples of individual instrument comparisons
AFTER
Serial # | RFNO3 | RIENH4 STD | RIESO4 STD | RIESO4 MIX | RIENH4 MIX
Original | 5.63e-12 | 5.11 | 0.76 | - | -
Calibrated | 3.21e-11 | 5.98 | 0.4 | 0.58 | 5.1
ACMCC: ACTRIS2_First intercomparison analysis and reporting

## Slide 23
The ACSM instrument is a research instrument,
The ACMCC is an excellent tool to improve our understanding of the ACSM response to different atmospheric compounds.
The data acquired during the ACMCC intercomparison exercises will be used to determine robust uncertainties for these instruments.
***ACTIONS***
Stations should provide calibration value RIE NH4, RF NO3 at least every 6 months
Encourage users to perform calibrations at several particle diameters to characterise if the TE changes before/after shipping instrument.
ACMCC: ACTRIS2_First intercomparison analysis and reporting